## Slide 1
MSc THESIS DEFENSE
Design of Piezoelectric Vibroacoustic
MEMS Sensors for Fully Implantable
Cochlear Implants
A tonotopic ScAlN bimorph cantilever array for passive, DSP-free frequency
separation across the speech intelligibility band
Ismail Musa Hamid
BME Faculty of Electrical Engineering and Informatics — Dept. of Electron Devices
Supervisors: Andras Pappe, DSc (BME VIK EET) • János Volk, DSc (HUN-REN EK-MFA)
Budapest, 2026

## Slide 2
01 — MOTIVATION
Hearing loss is a growing global burden
2.5B
people affected by hearing
loss projected by 2050
700M+
will require dedicated
hearing intervention
#3
leading global cause of
years lived with disability
Current cochlear implants all rely on externally worn hardware
External processor imposes bulk, battery maintenance, and moisture sensitivity
Adult CI recipients wear processors only ~10 hrs/day on average — well below clinical recommendation
87% remove devices during exercise or heat; 37% frequently forget to wear them entirely
Piezoelectric Vibroacoustic MEMS Sensors for FICI
2

## Slide 3
01 — MOTIVATION
The vision: a fully implantable cochlear implant
CONVENTIONAL CI — requires external hardware
Sound
External
Microphone
DSP
Processor
RF
Transmitter
Implanted
Receiver
Cochlear
Array
External processor — worn behind the ear — DSP consumes major share of power budget
PROPOSED FICI — fully implanted, passive, DSP-free
Sound
Ossicular
Vibration
MEMS Piezo
Cantilever Array
Implanted
Electronics
Cochlear
Array
All components subcutaneous — passive mechanical frequency separation — no digital signal processing
≤ 5×5 mm footprint and < 25 mg mass — middle-ear cavity constraint (~1 cm³)
Piezoelectric sensing: zero bias power, array of resonant elements gives passive channel separation
Piezoelectric Vibroacoustic MEMS Sensors for FICI
3

## Slide 4
01 — MOTIVATION
Thesis objectives
Benchmark modeling
Validated COMSOL FEM model of the commercial Vesper VA-1200 MEMS accelerometer
Experimental characterization
Scanning laser vibrometry of cantilever elements, correlated against FEM
Mechanical frequency tuning
FIB-based tuning to realize a cochlea-inspired, speech-range filter bank
Novel sensor design
Parametric FEM optimization of a piezoelectric bimorph cantilever array
Fabrication planning
MEMS process flow definition and photomask layout for cleanroom build
Cleanroom participation
Front-end processing (AlN sputtering, patterning/etching) and documentation
Piezoelectric Vibroacoustic MEMS Sensors for FICI
4

## Slide 5
02 — BIOLOGICAL INSPIRATION
Tonotopy: the cochlea's passive frequency map
base
apex → low f
base → high f
Basilar membrane: narrow & stiff near the base, wide & compliant toward the apex
A travelling wave peaks at a location set by stimulus frequency — high frequencies near the base, low frequencies near the apex
~35 mm membrane length resolves ~3,500 discrete frequency channels, spanning 20 Hz – 20 kHz
Frequency decomposition is purely mechanical — no active computation required
Design principle carried into Ch. 5: replace the single biological membrane with a discrete array of MEMS cantilevers, each geometrically tuned to a distinct frequency band.
Piezoelectric Vibroacoustic MEMS Sensors for FICI
5

## Slide 6
02 — BIOLOGICAL INSPIRATION
Why piezoelectric MEMS, and why AlN / ScAlN
Direct effect: mechanical strain → electric charge; zero bias power required
Cantilever arrays give passive mechanical frequency separation — array output is already channel-separated
AlN/ScAlN: CMOS-compatible, no poling step, sputter-deposited below 400°C
Sc alloying boosts d₃₁ / e₃₁ substantially over plain AlN, without sacrificing process compatibility
Material | CMOS-compatible | Poling needed
AlN / ScAlN | Yes | No
PZT | Challenging (Pb, 600–700°C) | Yes
PVDF | No (polymer) | Yes
Unimorph vs. Bimorph cross-section
Top electrode
Piezo layer (AlN)
Elastic substrate (Si)
Unimorph
Neutral axis
Bimorph (selected) — layers poled oppositely, outputs add constructively
Piezoelectric Vibroacoustic MEMS Sensors for FICI
6

## Slide 7
03 — BENCHMARK MODELING
Benchmark device: the Vesper VA-1200
Commercial AlN piezoelectric MEMS voice accelerometer, 2.9 × 2.76 × 0.9 mm package
Twelve-element cantilever array connected in series (SEM-verified)
Bimorph stack: two ScAlN films (~0.5 µm each) between three Mo electrodes
Geometry reconstructed from SEM tracing → imported into COMSOL as extruded solid
Sensitivity | −45 dBV/g @ 250 Hz
Resonant peak | 2.8 kHz (datasheet)
Noise floor | 1.5 mg rms, A-weighted
Supply current | 160 µA @ 1.8 V
Reconstructed cantilever geometry
670 µm
total cantilever length
1.15 µm
total bimorph stack thickness
ScAlN, x=0.4
Sc concentration used in FEM
Mo / Al
electrode / proof mass material
Physical characterization: package decapsulation → SEM imaging → EDS composition analysis → profile tracing → COMSOL solid model
Piezoelectric Vibroacoustic MEMS Sensors for FICI
7

## Slide 8
03 — BENCHMARK MODELING
COMSOL FEM validated against experiment
Chart: Frequency (kHz)
| 1st flexural | 2nd flexural
COMSOL simulated (kHz) | 3.1 | 30.44
Vibrometer measured (kHz) | 3.18 | 31.59
2.6%
1st mode deviation
3.10 vs 3.18 kHz
3.6%
2nd mode deviation
30.44 vs 31.59 kHz
Plain AlN under-predicted sensitivity (0.17 mV/g); Sc₀.₄Al₀.₆N adopted → 5.6 mV/g, matching datasheet trend
Series-connected 12-element array: simulated 5.28 mV/g → −45.55 dBV/g vs. −45 dBV/g datasheet
Model validated to within ~3.5% — used as the design tool for all subsequent optimization
Piezoelectric Vibroacoustic MEMS Sensors for FICI
8

## Slide 9
04 — POST-FABRICATION TUNING
FIB milling as a mechanical frequency-tuning tool
A single 140 µm longitudinal slit was FIB-milled into one VA-1200 cantilever (Design 7)
Slit reduces cross-sectional 2nd moment of area → lowers bending stiffness without materially changing modal mass
Verified with scanning laser vibrometry before / after milling
Demonstrates spatial selectivity: modified element resolves spectrally from untouched neighbours — the filterbank principle in practice
Chart: 140 µm slit, 29 µm from frame
| 1st mode | 2nd mode
Before milling (kHz) | 3.18 | 31.59
After milling (kHz) | 2.654 | 25.949
−16.5%
1st mode shift
3.18 → 2.654 kHz
4.3%
FEM vs. measured
Design 7 deviation
2.07–3.49
kHz design span
6 geometric variants simulated
Piezoelectric Vibroacoustic MEMS Sensors for FICI
9

## Slide 10
05 — NOVEL SENSOR DESIGN
Eight-element tonotopic bimorph filterbank
Each element is a pure ScAlN bimorph cantilever, resonance set by beam length and Si proof-mass geometry — operating below resonance in its flat, linear response region.
Group A — low frequency (fabricated)
A1 0.54 kHz
A2 1.06 kHz
A3 1.75 kHz
A4 3.53 kHz
150 µm beam + variable 150–450 µm Si proof mass (A4: 700 µm beam, no mass)
Group C — high frequency (design)
C1 5.35 kHz
C2 7.11 kHz
C3 9.25 kHz
C4 14.90 kHz
160–340 µm short beams with thin 75 µm proof mass for compact high-f channels
Common layer stack: 0.5 µm AlN, 300 µm beam width across all 8 elements
Frequency ordering set by geometric ratios — robust to absolute ScAlN material uncertainty
Piezoelectric Vibroacoustic MEMS Sensors for FICI
10

## Slide 11
05 — NOVEL SENSOR DESIGN
Analytical model precedes the FEM sweep
Composite Euler-Bernoulli beam theory (Erturk & Inman, 2008; Rao, 2019) for the symmetric ScAlN bimorph:
k_eff = 2·w·E_p·t_p³ / L³
Tip stiffness — cubic in beam length
f₁ = (1/2π)·√(k_eff / m_eff)
Fundamental resonant frequency
S_v,A ∝ M_tip · L / (w · t_p)
Sensitivity, elements WITH proof mass
S_v,B ∝ L² (independent of w, t_p)
Sensitivity, elements WITHOUT proof mass
Key insight: for proof-mass-free elements, adding piezoelectric volume increases both charge AND stiffness proportionally — sensitivity depends only on beam length. Material e₃₁,f/E ratio, not cross-section, is the lever for these channels.
Piezoelectric Vibroacoustic MEMS Sensors for FICI
11

## Slide 12
06 — FILTERBANK PERFORMANCE
Filterbank spans the full speech intelligibility band
Chart: 
| A1 | A2 | A3 | A4 | C1 | C2 | C3 | C4
Resonant frequency f₁ (kHz) | 0.54 | 1.06 | 1.75 | 3.53 | 5.35 | 7.11 | 9.25 | 14.9
0.29–8.05
kHz usable bandwidth
sub-resonance, f₁/1.85
0.54–14.90
kHz resonance span
across all 8 elements
300–8000
Hz published target
band substantially met
Piezoelectric Vibroacoustic MEMS Sensors for FICI
12

## Slide 13
06 — FILTERBANK PERFORMANCE
Noise-equivalent acceleration vs. published target
Chart: Target: < 10 µg/√Hz at 1 kHz
| A1 | A2 | A3 | A4 | C1 | C2 | C3 | C4
NEA (µg/√Hz) | 1.29 | 2.69 | 4.84 | 82.2 | 32.5 | 39.6 | 64 | 237
1.27–3.92
µg/√Hz for A1–A3
2–8× below the 10 µg/√Hz target
20
measured quality factor Q
from VA-1200 vibrometry, air damping
A4–C4
electronic-noise-limited
readout ASIC is the design lever here
Piezoelectric Vibroacoustic MEMS Sensors for FICI
13

## Slide 14
06 — FILTERBANK PERFORMANCE
Compliance with published design targets
Metric | Target | Achieved | Status
Frequency range | 0.1 – 8 kHz | 0.29 – 8.05 kHz | Substantially met
Sensitivity | 1 – 10 mV/g | 0.048 – 9.98 mV/g | Met for A1–C2
Noise floor | < 10 µg/√Hz @ 1kHz | 1.27 – 3.92 µg/√Hz (A1–A3) | Met, lower band
Resonant freq. | 1 – 8 kHz multiband | 0.54 – 14.90 kHz, 8 ch. | Met
Device mass | sub-mg proof masses | 0.078 – 47.23 µg | Met
Die footprint | < 1 mm² sensing | ≈ 2.03 mm² active area | Partially met
Group C sensitivity/noise shortfall reflects the fundamental sensitivity–bandwidth trade-off, not a limitation of the AlN bimorph mechanism itself.
Piezoelectric Vibroacoustic MEMS Sensors for FICI
14

## Slide 15
07 — FABRICATION
Twelve-step process, seven photomasks
1
Substrate prep +
thermal SiO₂
2
Bottom Mo
electrode (M1)
3
AlN layer 1
sputter
4
Middle Mo
electrode (M2)
5
AlN layer 2
sputter
6
Top Mo
electrode (M3)
7
Frontside AlN
wet etch (M4)
8
Au contact
pads (M5)
9
Backside Al
hard mask (M6)
10
Backside PR
mask (M7)
11
Nested Bosch
DRIE (2-depth)
12
Final HF
release
380 µm
DSP Si wafer
500 nm thermal SiO₂ etch stop
Mo / ScAlN
electrodes / piezo
reactive RF magnetron sputter
10 µm
min. feature size
5× above contact-aligner limit
Piezoelectric Vibroacoustic MEMS Sensors for FICI
15

## Slide 16
07 — FABRICATION
Photomask layout: 3.5 × 2.5 mm sensor die
Group A (long beams)
Group C (short beams)
Release
cavity
3.5 mm × 2.5 mm die
Designed in CleWin, minimum feature size 10 µm (5× process margin)
3-pad electrode column per cantilever, 250 µm pitch, nested Au/via/Mo pad stack
100 µm cleavage flange for controlled die separation after HF release
IR transmission alignment for frontside-to-backside overlay (±5 µm budget)
Status at time of writing: mask set, process flow, and deposition qualification complete; device wafer fabrication planned as next step.
Piezoelectric Vibroacoustic MEMS Sensors for FICI
16

## Slide 17
08 — SUMMARY
Summary of contributions
Validated FEM framework
Benchmarked against VA-1200: 2.6% / 3.6% deviation on 1st / 2nd flexural modes
FIB frequency-trimming method
Demonstrated 16.5% deterministic downshift via a single milled slit
Closed-form analytical model
Bimorph stiffness, mass, sensitivity relations — S ∝ L³, f₁ ∝ 1/L² trade-off
8-element tonotopic filterbank
0.54–14.90 kHz resonance; NEA 1.29–4.84 µg/√Hz for the lower band
Complete fabrication plan
12-step / 7-mask process defined and qualified for HUN-REN EK-MFA
Piezoelectric Vibroacoustic MEMS Sensors for FICI
17

## Slide 18
08 — SUMMARY
Limitations and further development
LIMITATIONS
FEM used AlN, not ScAlN, material properties (deposition not yet characterized)
Group C element sensitivity falls below the 1 mV/g target
Device wafer fabrication has not yet commenced
NEXT STEPS
Execute the 12-step wafer run; verify resonances by vibrometry vs. FEM (±5% target)
Measure e₃₁,f of sputtered ScAlN on calibration beams to remove the largest model uncertainty
Low-noise charge-amplifier ASIC (< 1 µV/√Hz) to close the Group C sensitivity gap
FIB trimming look-up table for individual post-fabrication element correction
Vacuum packaging to raise Q by 3–5× and suppress thermomechanical noise
System-level path: readout ASIC (180 nm CMOS) + intracochlear stimulator; in-vivo validation in a deafened animal model before clinical translation.
Piezoelectric Vibroacoustic MEMS Sensors for FICI
18

## Slide 19
Thank you
Questions & Discussion
Ismail Musa Hamid — BME VIK EET × HUN-REN EK-MFA